+2-011-23849195

EEV 20.1

ETNAFS@GMail.com

USER MANUAL

ETNAFS EMERGENCY VENTILATOR

## EEV 20 ventilator

### Parameters

- range 5 to 37 cm H2O adjustable in 1 cmH2O increment(+-0.001cmH2O)

- Monitor wirelessly over Wi-Fi network

- Air/Oxygen supply shouldn’t exceed 20 PSI

- Expiratory ratio (I:E) adjustable in the range 1.0 to 4.0 in 1 increment

- range 10 to 30 breaths per minute in increments of 1 (+-1breath)

- Tidal Volume (Vt) optionally Range 250 to 800 ml in steps of 10ml (+-30ml)

- 1 Hour rechargeable battery backup if power failure

- Oxygen percent control range 21% to 100% increments of 1% (+-2 %)

- The ability to input body weight and have volume calculated as, for example, 6ml/kg of ideal body weight

## Alarms

- gas or electricity supply failure

- machine switched off while in mandatory ventilation mode

- inspiratory airway pressure exceeded

- inspiratory pressure not achieved (equivalent to disconnection alarm)

- tidal volume not achieved or exceeded

# Monitoring

### Modes

  - Continuous Positive Airway Pressure (CPAP)

  - Continuous Mandatory Ventilation (CMV)

  - Synchronous Intermittent Mandatory Ventilation (SIMV)

  - Airway Pressure Release Ventilation (APRV)

  - Pressure Support Ventilation (PSV)

# Wireless monitoring

- Download ETNAFS android app from app store

- Scan the device QR code

- Monitor the device from your phone

# Device control

# Choose the parameter you want to change

# Write the input

# To change the mode just press on the mode button

# To change the I:E ratio press on the I:E button